## Comp_Sht
Uncertainty Meaurement for Thermometer/Sensor
1) Uncertainty due to Reference Standard Used Un/ U1 Standard Uncertainty of SPRT and Read-out/ Indicator)
Standard Platinum Resistance Thermometer
| Description |  |  |  |  | Digital Thermometer with PT100 Probe
[6 empty] | Confidence level = 95 % |  |  |  | Standard k factor (from cal cert) k= |  | 2
[6 empty] | Degrees of freedom (infinite), df1= |  |  | 1e+26 | Sensitivity Coefficient, sc1= |  | 1
| Actual Temperature °C |  |  | Digital Thermometer Indicated Values °C |  |  | Corrections |  | Uncertainty (°C) |  | Uncertainty Un
[9 empty] | Standard uncertainty (from cal cert.), Us |  | °C |  | UN | ∑UNi
| 0 | [5 empty] | 0 |  | 0.023 |  | 0.0115 |  |  | k
| 50 | [5 empty] | 50 |  | 0.023 |  | 0.0115
| 100 | [5 empty] | 100 |  | 0.023 |  | 0.0115
Environment Condition
[14 empty] | Time | Temperature (oC) | Humidity (% RH)
2) Uncertainty due to Repeatability (Ur/ U2) | [12 empty] | Start | 10:00 | 24.6 | 53.8
| x | 36 | 309.15 | 100 | 121 |  | No. Of Trials (n)= |  |  | 3 |  |  | End | 14:30 | 24.3 | 54.2
|  | Testpoint 1 |  | Testpoint 2 | Testpoint 3 |  | Degrees of freedom (n-1), df2= |  |  | 2 |  |  |  | Average | 24.45 | 54
| Trial 1 |  | 273.15 |  |  | °C | Sensitivity Coefficient, sc2= |  |  | 1
| Trial 2 |  | 273.15 |  |  | °C
| Trial 3 |  | 273.15 |  |  | °C | [5 empty] | Standard dev/ Sr = √ (1/n-1)[∑(x-x')2]
| Ave, x' | #DIV/0! | 273.15 | #DIV/0! | #DIV/0! |  | ∑(x-x')2 | #DIV/0! | #DIV/0! | #DIV/0! |  | Ur = | Sr
x-x' | T1 | #DIV/0! | 0 | #DIV/0! | #DIV/0! |  | Sr = | #DIV/0! | #DIV/0! | #DIV/0! |  |  | √n
| T2 | #DIV/0! | 0 | #DIV/0! | #DIV/0! |  | LONG CUT STANDARD DEVIATION | [6 empty] | °C
| T3 | #DIV/0! | 0 | #DIV/0! | #DIV/0! | [7 empty] | Testpoint 1 | #DIV/0!
'(x-x')2 | T1 | #DIV/0! |  | #DIV/0! | #DIV/0! | [7 empty] | Testpoint 2 | #DIV/0!
| T2 | #DIV/0! |  | #DIV/0! | #DIV/0! | [7 empty] | Testpoint 3 | #DIV/0!
| T3 | #DIV/0! |  | #DIV/0! | #DIV/0! |  | Sr (excel) | #DIV/0! | #DIV/0! | #DIV/0!
3) Uncertainty due to readability (Ud/ U3)
| Resolution/Graduation, Rg |  |  |  | °C
| Readability, Rd |  |  |  | (0.1 for 1/10, 0.2 for 1/5, 0.5 for 1/2, 1 for 1/1)
| Ud= (Rg* Rd )/ √3 |  |  | 0
| Absolute uncertainty of device |  |  | 0.5
| Measured value |  |  | 25.2
| Relative uncertainty %R |  |  | 1.9841
|  |  |  | 5 | (5 best, 10 good, 20 poor)
| Degrees of freedom df3 |  |  | 200
1/2(100/%R)^2 = 0.5(100/%R)^2
| Sensitivity Coefficient, sc3 |  |  | 1

## Testpoints
For 1ST testpoint
Component | units | Distribution | Type | dfi | ui | sci | uisci | (uisci)2 | (uisci)4/dfi
Un: Std Uncertainty | °C | Normal | B | 1e+26 | 0.0115 | 1 | 0.0115 | 0.00013225 | 1.749e-34
Ur: Repeatability | °C | Normal | A | 2 | #DIV/0! | 1 | #DIV/0! | #DIV/0! | #DIV/0!
Ud: Readability/Resolution | °C | Rect. | B | 200 | 0 | 1 | 0 | 0 | 0
Sum | [7 empty] | #DIV/0! | #DIV/0!
Combined Standard uncertainty = square root of the sum of (uisci)2, Uc | [7 empty] | #DIV/0! | °C
Effective degrees of freedom = (Uc)4 / sum of (uisci)4/dfi | [6 empty] | Veff | #DIV/0!
Coverage Factor = Students (for Veff and C.L. 95%), K | [7 empty] | #DIV/0!
Expanded Uncertainty @ 95% C.L., Ue | [7 empty] | #DIV/0! | °C
For 2ND testpoint
Component | units | Distribution | Type | dfi | ui | sci | uisci | (uisci)2 | (uisci)4/dfi
Un: Std Uncertainty | °C | Normal | B | 1e+26 | 0.0115 | 1 | 0.0115 | 0.00013225 | 1.749e-34
Ur: Repeatability | °C | Normal | A | 2 | #DIV/0! | 1 | #DIV/0! | #DIV/0! | #DIV/0!
Ud: Readability/Resolution | °C | Rect. | B | 200 | 0 | 1 | 0 | 0 | 0
Sum | [7 empty] | #DIV/0! | #DIV/0!
Combined Standard uncertainty = square root of the sum of (uisci)2, Uc | [7 empty] | #DIV/0! | °C
Effective degrees of freedom = (Uc)4 / sum of (uisci)4/dfi | [6 empty] | Veff | #DIV/0!
Coverage Factor = Students (for Veff and C.L. 95%), K | [7 empty] | #DIV/0!
Expanded Uncertainty @ 95% C.L., Ue | [7 empty] | #DIV/0! | °C
For 3RD testpoint
Component | units | Distribution | Type | dfi | ui | sci | uisci | (uisci)2 | (uisci)4/dfi
Un: Std Uncertainty | °C | Normal | B | 1e+26 | 0.0115 | 1 | 0.0115 | 0.00013225 | 1.749e-34
Ur: Repeatability | °C | Normal | A | 2 | #DIV/0! | 1 | #DIV/0! | #DIV/0! | #DIV/0!
Ud: Readability/Resolution | °C | Rect. | B | 200 | 0 | 1 | 0 | 0 | 0
Sum | [7 empty] | #DIV/0! | #DIV/0!
Combined Standard uncertainty = square root of the sum of (uisci)2, Uc | [7 empty] | #DIV/0! | °C
Effective degrees of freedom = (Uc)4 / sum of (uisci)4/dfi | [6 empty] | Veff | #DIV/0!
Coverage Factor = Students (for Veff and C.L. 95%), K | [7 empty] | #DIV/0!
Expanded Uncertainty @ 95% C.L., Ue | [7 empty] | #DIV/0! | °C

## Cal. Cert.
REGIONAL STANDARDS AND TESTING LABORATORY
CALIBRATION CERTIFICATE
| Reference No. |  | :
| Sample No. |  | :
| Date Submitted |  | :
| Date Calibrated |  | :
| Calibration Item |  | :
|  | Make | :
|  | Model No. | :
|  | Serial No. | :
|  | Resolution | : | Readabilty |  | :
|  | Customer | :
|  | Address | :
MEASUREMENT RESULTS:
This certifies that the above THERMOMETER/ SENSOR was calibrated by the Regional Metrology Laboratory at Government Center, Sevilla, City of San Fernando, La Union under the following environmental conditions:
| Ambient Temperature (º C) |  | : | º C
| Relative Humidity (%) |  | : | %
| Standard Reading º C |  | UUT Reading º C |  | Correction º C |  | Uncertainty of Measurement
| 0 |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
| 0 |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
| 0 |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
The procedure is in accordance with the laboratory’s in-house calibration method (CM-MTR-009) with the following results at the indicated test points. The specified expanded uncertainty is at a confidence level of approximately 95% with a coverage factor of k=2.
STANDARDS USED AND TRACEABILITY
| Name of Standard |  | Calibration Certificate No. |  |  | Issuing Lab/Traceability
| Temperature Readout Isotech TTI-10 |  | 102023-DT-0089 |  |  | NML-Philippines, ITDI
| Standard Platinum Resistance Thermometer |  | 102023-DT-0089
REMARKS:
| • | The results given in this report are obtained at the time of the test and refer only to the particular instrument submitted. This report shall not be reproduced except in full, without the written approval of the laboratory.
| • | This instrument was calibrated using reference standard traceable to SI Units of measurement through National Metrology Laboratory
| • | The End user should determine the suitability of equipment for it's intended use.
| • | No adjustments were performed on the thermometer
JULIUS R. ALVIOR
Calibration Engineer
BERNADINE P. SUNIEGA
Technical Manager
Postal Address: Government Center, Sevilla Tel. No.: (072) 242-0663 City of San Fernando, La Union Mobile No. : +63 9684435399
e-mail address: rml@region1.dost.gov.ph URL: http://region1.dost.gov.ph
Postal Address: DMMMSU-MLU Campus, P. O. Box 117 Tel. No.: (072) 242-0663 – LU, (077) 722-3347 - IS
City of San Fernando, 2500 La Union Fax. No.: (072) 888-3399 - ASD
e-mail address: dostro1@dost.gov.ph URL: http://region1.dost.gov.ph
Republic of the Philippines
DEPARTMENT OF SCIENCE AND TECHNOLOGY
Regional Office No. I
Republic of the Philippines
DEPARTMENT OF SCIENCE AND TECHNOLOGY
Regional Office No. I
Page 1 of 2
Page 2 of 2