## WORKSHEET
Regional Standard and Testing Laboratory - I
Regional Metrology Laboratory
Data Sheet for Testing of Sphygmomanometer
| Customer: |  |  |  |  | Date submitted:
| Address: |  |  |  |  | Calibration Date:
| Request Reference No. |  |  |  |  | Calibrated by
| Sample no. |  |  |  |  | Place of Calibration:
Identity and Specifications of Sphygmomanometer
| Model |  |  |  | Range
| Serial No. |  |  |  | Accuracy
Environment Condition
|  | Time |  | Temperature (oC) | Humidity (% RH) | Pressure (hPa)
Start
End
-
Standard Used
|  | Type | : | Digital Pressure Calibrator
|  | Model/ Maker | : | ADT672-05-GP15-BAR-N
|  | Measurement Range | : | 0 -1 bar (0 – 750 mmHg)
|  | Accuracy (% FS) | : | 0.05
|  | Serial No. | : | 2731706006
|  | Certificate No. | : | 17101709
|  | Traceability | : | National Institute of Standards and Technology (NIST)
|  | Calibration Date | : | October 17, 2017
Test for the Maximum Permissible Errors of the Cuff Pressure Indication
STANDARD READINGS (IPRT)
| APPLIED PRESSURE (mmHg) | X1 | X2 | X3 | X4 | MEAN
|  | up | down | up | down | IPRT
| 0 | 0 | 0 | 0 | 0
50
100
150
200
250
300
UNIT UNDER TEST READINGS (UUT)
| APPLIED PRESSURE (mmHg) | X1 | X2 | X3 | X4 | MEAN
|  | up | down | up | down | up | down | UUT
| 0 | 0 | 0 | 0 | 0
| 50 | 50 | 50 | 50 | 50
| 100 | 100 | 100 | 100 | 100
| 150 | 150 | 150 | 150 | 150
| 200 | 200 | 200 | 200 | 200
| 250 | 250 | 250 | 250 | 250
| 300 | 300 | 300 | 300 | 300
Test for Air Leakage of the Pneumatic System
| Applied Pressure (mmHg) |  | Sample Reading (mmHg) |  |  |  | Rate of Pressure Loss
|  |  | 1st Reading |  | After 5 minutes
| 60 |  | 60 |  | 60
| 120 |  | 120 |  | 120
| 180 |  | 180 |  | 180
| 240 |  | 240 |  | 238
| 300 |  | 300 |  | 296
Test for the Rapid Exhaust Valve
| Applied Pressure (mmHg) |  |  | Time for the Pressure Reduction to reach ≤ 15 mmHg

## Data Sheet
Customer | : |  |  |  |  | Particulars | ;
Address | : |  |  |  |  | Model | :
Reference No. | : |  |  |  |  | Serial No. | :
Sample No. | : |  |  |  |  | Range | :
Date Submitted | : |  |  |  |  | Graduation | :
Date Calibrated | :
Place of Calibration | :
Environment Condition
|  | Time | Temperature (oC) | Humidity (% RH)
Start
End
|  | Average | #DIV/0! | #DIV/0!
Uncertainty Calculation (DKD R-6-1 Calibration of Pressure Gauges) | [7 empty] | 1 kPa= | 7.5006 | mmHg | 1 bar = | 750.0616 | mmHg | 100 | kPa
STANDARD READINGS (IPRT)
APPLIED PRESSURE (mmHg) | X1 | X2 | X3 | X4 | MEAN |  |  | APPLIED (kPa) | X1 | X2 | X3 | X4 | MEAN
| up | down | up | down | IPRT |  |  |  | up | down | up | down | IPRT
0 | 0 | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0
50 |  |  |  |  | #DIV/0! |  |  | 6.6661 | 0 | 0 | 0 | 0 | 0
100 |  |  |  |  | #DIV/0! |  |  | 13.3322 | 0 | 0 | 0 | 0 | 0
150 |  |  |  |  | #DIV/0! |  |  | 19.9983 | 0 | 0 | 0 | 0 | 0
200 |  |  |  |  | #DIV/0! |  |  | 26.6644 | 0 | 0 | 0 | 0 | 0
250 |  |  |  |  | #DIV/0! |  |  | 33.3305 | 0 | 0 | 0 | 0 | 0
300 |  |  |  |  | #DIV/0! |  |  | 39.9967 | 0 | 0 | 0 | 0 | 0
UNIT UNDER TEST READINGS (UUT)
APPLIED PRESSURE (mmHg) | X1 | X2 | X3 | X4 | MEAN |  |  | APPLIED (kPa) | X1 | X2 | X3 | X4 | MEAN
| up | down | up | down | up | down | UUT |  | up | down | up | down | up | down | UUT
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 6.6661 | 6.6661 | 6.6661 | 6.6661 | 6.6661 | 6.6661 | 6.6661 | 6.6661
100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 13.3322 | 13.3322 | 13.3322 | 13.3322 | 13.3322 | 13.3322 | 13.3322 | 13.3322
150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 19.9983 | 19.9983 | 19.9983 | 19.9983 | 19.9983 | 19.9983 | 19.9983 | 19.9983
200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 26.6644 | 26.6644 | 26.6644 | 26.6644 | 26.6644 | 26.6644 | 26.6644 | 26.6644
250 | 250 | 250 | 250 | 250 | 250 | 250 | 250 | 33.3305 | 33.3305 | 33.3305 | 33.3305 | 33.3305 | 33.3305 | 33.3305 | 33.3305
300 | 300 | 300 | 300 | 300 | 300 | 300 | 300 | 39.9967 | 39.9967 | 39.9967 | 39.9967 | 39.9967 | 39.9967 | 39.9967 | 39.9967
DEVIATION |  |  | Max hysteresis Error (mmHg)
mean (UUT) - mean (IPRT) |  |  | X1-X2 | X3-X4
mmHg | kPa |  | 0 | 0
0 | 0 |  | 0 | 0
#DIV/0! | 6.6661 |  | 0 | 0
#DIV/0! | 13.3322 |  | 0 | 0
#DIV/0! | 19.9983 |  | 0 | 0
#DIV/0! | 26.6644 |  | 0 | 0
#DIV/0! | 33.3305 |  | 0 | 0
#DIV/0! | 39.9967 |  | 0 | 0
Applied Pressure (mmHg) |  | Sample Reading (mmHg) |  |  |  | Rate of Pressure Loss
|  | 1st Reading |  | After 5 minutes
|  | 60 |  | 60 |  | 0
|  | 120 |  | 120 |  | 0
|  | 180 |  | 180 |  | 0
|  | 240 |  | 238 |  | 0.4
|  | 300 |  | 296 |  | 0.8
Applied Pressure (mmHg) |  |  | Time for the Pressure Reduction to reach ≤ 15 mmHg
300 |  |  |  |  | 5.1

## Calibration Report
CALIBRATION CERTIFICATE
Customer |  | : | 0 |  |  |  | Particulars | ;
Address |  | : | 0 |  |  |  | Model | : | 0
Reference No. |  | : | 0 |  |  |  | Serial No. | : | 0
Sample No. |  | : | 0 |  |  |  | Range | : | 0
Date Submitted |  | : | 0 |  |  |  | Graduation | : | 0
Date Calibrated |  | : | 0
Place of Calibration |  | : | 0
TEST RESULTS:
This certifies that the above scale has been calibrated and adjusted to its nominal capacity by this laboratory by performing the following tests;
I. Test for the Maximum Permissible Errors of the Cuff Pressure Indication
Applied Pressure (mmHg) |  | Sample Reading (mmHg) |  |  |  | Maximum Deviation | Maximum Permissible Error
|  | Increasing |  | Decreasing
0 |  | 0 |  | 0 |  | 0 | within ± 4 mmHg
|  | 50 |  | 50 |  | #DIV/0! | within ± 4 mmHg
|  | 100 |  | 100 |  | #DIV/0! | within ± 4 mmHg
|  | 150 |  | 150 |  | #DIV/0! | within ± 4 mmHg
|  | 200 |  | 200 |  | #DIV/0! | within ± 4 mmHg
|  | 250 |  | 250 |  | #DIV/0! | within ± 4 mmHg
|  | 300 |  | 300 |  | #DIV/0! | within ± 4 mmHg
II. Test for Hysteresis Error
Maximum Hysteresis Error |  |  |  |  | Maximum Permissible Error
0 |  |  |  |  | within ± 4 mmHg
III. Test for Air Leakage of the Pneumatic System
Applied Pressure (mmHg) |  | Sample Reading (mmHg) |  |  |  | Rate of Pressure Loss | Maximum Permissible Error
|  | 1st Reading |  | After 5 minutes
|  | 60 |  | 60 |  | 0 | ≤4.0 mmHg/min
|  | 120 |  | 120 |  | 0 | ≤4.0 mmHg/min
|  | 180 |  | 180 |  | 0 | ≤4.0 mmHg/min
|  | 240 |  | 238 |  | 0.4 | ≤4.0 mmHg/min
|  | 300 |  | 296 |  | 0.8 | ≤4.0 mmHg/min
IV. Test for the Rapid Exhaust Valve
Applied Pressure (mmHg) |  |  |  | Time for the Pressure Reduction to reach ≤ 15 mmHg |  |  | Maximum Permissible Error
300 |  |  |  | 5.1 |  |  | < 10 seconds
This certifies that the above Non - invasive Sphygmomanometer was Tested by the Regional Metrology Laboratory at Government Center, Sevilla, City of San Fernando, La Union.
The sphygmomanometer was tested according to its normal use with the pressure inlet connection at the bottom. The sphygmomanometer has been tested from 0 mmHg to 300 mmHg positive gauge pressures under the following environmental conditions:
| Ambient Temperature (°C) |  |  | : | #DIV/0! | °C
| Relative Humidity (RH) |  |  | : | #DIV/0! | %
The sphygmomanometer was subjected to air leakage, rapid exhaust valve, cuff pressure indication, and hysteresis tests in accordance with the laboratory’s in-house calibration method CM-MET-008 “Performance Testing of Sphygmomanometer” based on the International Recommendation OIML R 16-1 Edition 2002 “Non-invasive Mechanical Sphygmomanometers”.
STANDARD USED AND TRACEABILITY:
Standard Gauge
Type |  | : | Digital Pressure Calibrator
Model/ Maker |  | : | ADT672-05-GP15-BAR-N
Measurement Range |  | : | 0 -1 bar (0 – 750 mmHg)
Accuracy (% FS) |  | : | 0.05
Serial No. |  | : | 2731706006
Certificate No. |  | : | 102023-INS-08090
Traceability |  | : | Metal Industry Research and Development Center
Calibration Date |  | : | 2025-10-10
REMARKS:
• The general condition and workmanship of the instrument are SATISFACTORY
• The results given in this report were obtained at the time of the test and refer only to the particular sphygmomanometer submitted under the conditions indicated. The certificate is not intended to be representative of similar items. If the unit is modified or damaged in anyway, the results may be rendered invalid and will require re-verification.
• The user should determine suitability of the sphygmomanometer for its intended use.
• No adjustments were performed on the sphygmomanometer.
• The sphygmomanometer reading may be corrected by applying a correction of 1 mmHg. (e.g. Sphygmomanometer Reading: 120/80; After correction: 121/81, BASED ON THE CUFF PRESSURE INDICATION)
• The user is obliged to have the unit recalibrated at appropriate intervals.
• This report shall not be reproduced except in full, without the written approval of the laboratory.
Calibrated by: | [5 empty] | Certified by:
MA. FERNANDA I. BANDA | [5 empty] | BERNADINE P. SUNIEGA
Calibration Engineer |  |  |  |  | Technical Manager
Republic of the Philippines
DEPARTMENT OF SCIENCE AND TECHNOLOGY
Regional Office No. I
REGIONAL STANDARDS AND TESTING LABORATORY
Postal Address: DMMMSU-MLU Campus, P. O. Box 117 Tel. No.: (072) 242-0663 – LU, (077) 722-3347 - IS
City of San Fernando, 2500 La Union Fax. No.: (072) 888-3399 - ASD
e-mail address: dostro1@dost.gov.ph URL: http://region1.dost.gov.ph
Page 1 of 2
Republic of the Philippines
DEPARTMENT OF SCIENCE AND TECHNOLOGY
Regional Office No. I
REGIONAL STANDARDS AND TESTING LABORATORY
Postal Address: DMMMSU-MLU Campus, P. O. Box 117 Tel. No.: (072) 242-0663 – LU, (077) 722-3347 - IS
City of San Fernando, 2500 La Union Fax. No.: (072) 888-3399 - ASD
e-mail address: dostro1@dost.gov.ph URL: http://region1.dost.gov.ph
Page 2 of 2